## Slide 1: AIR-AD: EVS Concept Notes

## Slide 2: Rapidly manufactured ventilator system specification:
https://www.gov.uk/government/publications/coronavirus-covid-19-ventilator-supply-specification/rapidly-manufactured-ventilator-system-specification
Notes: ICU min req specs:
1) volume control: you can set the tidal volume for each breath
2) set the respiratory rate
3 be able to set the PEEP, or at least a constant PEEP that is chosen by manufacturing team
Preferably able to do these:
A) set inspiratory flow
B) measure PIP (peak inspiratory pressure)

## Slide 3: Design Inputs Summary - Ventilation
General
Mode: volume or pressure control mode (Volume preferred)
Volume control: 250ml – 500ml (+/- 50ml) in at least 2 increments (ex 350 ml and 450 ml)
Breathing Rate: 10-40 Breaths per minute
Inspiratory/Expiratory (I:E) ratio fixed at 2 or adjustable 1.0 to 3.
Set PEEP: 5-20 cmH2O or one fixed (setting accuracy of PEEP (+/- 10%))
Maximum PIP 40-50 cmH2O
Optional:
Measure PIP (peak inspiratory pressure) 50 cmH2O
Flow, pressure, and volume waveforms
Notes: ICU min req specs:
1) volume control: you can set the tidal volume for each breath
2) set the respiratory rate
3 be able to set the PEEP, or at least a constant PEEP that is chosen by manufacturing team
Preferably able to do these:
A) set inspiratory flow
B) measure PIP (peak inspiratory pressure)

## Slide 4: Design Inputs Summary - Operation
The devices is valuable if it can reliably provide the clinical function, while not introducing substantial risk
Must interface with standard pneumatic connectors
Can’t interfere or damage other equipment
Must operate reliably 24 hours a day for at least 14 days.
Must be intuitive to use for medically trained professional

## Slide 5: Design Inputs Summary – Part Availability
Part or direct substitute must be available in quantities of 120
Device should use standard, industrial, certified components whenever possible
Device should be built using a modular architecture allowing for swapping of components based on availability

## Slide 6: Power Subsystem
AC PLUG
AC to 24 VDC
24V UPS
ACTUATOR DRIVER
CONTROLLER
BATTERY
ON/OFF
SWITCH
(Keep External)
SSR Latch
OFF SWITCH
(Optional)
DEVICE
OUTSIDE

## Slide 7: Pneumatic Subsystem
DEVICE
OUTSIDE
Ambubag/Cylinder
Pump
One-way Valve
22mm Air Port
Overpressure Valve 40 cmH2O
Pressure Port Barb
Overpressure Valve 60 cmH2O
(optional)
O2 Mixer
ACTUATOR
(electrically powered)
Emergency Air Input
Anti asphyxiation Valve
Air Input

## Slide 8: External Air Circuit
DEVICE
OUTSIDE
22mm Air Port
Pressure Port Barb
O2 Mixer
(find is available)
Patient Tubing
Mask/Tube
PEEP Valve
O2 Cylinder

## Slide 9: Control Subsystem:
Actuator Driver
Controller
(PLC unitronix)
Actuator
Position Feedback
Sensor
Endstop
Output Flow Sensor
Ambubag/Cylinder
Pump
Patient Pressure Sensor
Auxiliary Pressure Sensor
Tidal Vol Pot. (#1)
Breathing Rate Pot. (#2)
I/E Potentiometer (#3)
GUI Screen
(optional)
DEVICE
On-Demand Inhale Button
(optional)

## Slide 10: User Interface Concept
Tidal
Rate
I/E
Push On
AC Cord Wrap Handles
Handle
AC IN
Power Switch
Critical User Interface
Air and Pressure Port
Inlet Port & Emergency Inlet
Programmable LCD Display
Notes: ICU min req specs:
1) volume control: you can set the tidal volume for each breath
2) set the respiratory rate
3 be able to set the PEEP, or at least a constant PEEP that is chosen by manufacturing team
Preferably able to do these:
A) set inspiratory flow
B) measure PIP (peak inspiratory pressure)

## Slide 11: Piston Actuator Concept #1
Linear Actuator
Bellows
Plate
Cylinder
Check Valve #1
Check Valve #2
Stroke 3-5”
Stroke vs max reps vs freq
Flow rates can be linear
Physical limit switches
Some linear actuators have encoder
Plate sitting on 4 shafts with bearings
Complexity:
How to lead the actuator
Tolerance between the cylinder
Can be adressed with shafts
Notes: ICU min req specs:
1) volume control: you can set the tidal volume for each breath
2) set the respiratory rate
3 be able to set the PEEP, or at least a constant PEEP that is chosen by manufacturing team
Preferably able to do these:
A) set inspiratory flow
B) measure PIP (peak inspiratory pressure)

## Slide 12: Piston Actuator Concept #2
Linear Actuator
Plate
Cylinder
Check Valve #1
Check Valve #2
Stroke 3-5”
Floating O-Ring
Look into piston pumps
Existing components
Like cylinders
Notes: ICU min req specs:
1) volume control: you can set the tidal volume for each breath
2) set the respiratory rate
3 be able to set the PEEP, or at least a constant PEEP that is chosen by manufacturing team
Preferably able to do these:
A) set inspiratory flow
B) measure PIP (peak inspiratory pressure)

## Slide 13: Piston Actuator Concept #3
Linear Actuator
Bellows
Plate
Cylinder
Check Valve #1
Check Valve #2
Stroke 3-5”
Notes: ICU min req specs:
1) volume control: you can set the tidal volume for each breath
2) set the respiratory rate
3 be able to set the PEEP, or at least a constant PEEP that is chosen by manufacturing team
Preferably able to do these:
A) set inspiratory flow
B) measure PIP (peak inspiratory pressure)

## Slide 14: Piston Actuator Concept #4
Plate
Cylinder
Check Valve #1
Check Valve #2
O-Ring
Scissor type linkage
Notes: ICU min req specs:
1) volume control: you can set the tidal volume for each breath
2) set the respiratory rate
3 be able to set the PEEP, or at least a constant PEEP that is chosen by manufacturing team
Preferably able to do these:
A) set inspiratory flow
B) measure PIP (peak inspiratory pressure)

## Slide 15: Penn ICU Design Inputs :
Main:
Volume control – set tidal volume breathing
Set respiratory rate
PEEP set
Optional:
Set inspiratory flow
Measure PIP (peak inspiratory pressure)
Notes: ICU min req specs:
1) volume control: you can set the tidal volume for each breath
2) set the respiratory rate
3 be able to set the PEEP, or at least a constant PEEP that is chosen by manufacturing team
Preferably able to do these:
A) set inspiratory flow
B) measure PIP (peak inspiratory pressure)